## Document History
DOCUMENT HISTORY
Product name | Differential diagnostic PCR test (Pilot Test Case N°1)
Product reference
Date | 2021-04-14 | IP rights | Instituto de Medicina Molecular João Lobo Antunes (iMM)
Version # | Partner | Changes | Comments | Date
v0 | VITO | Document creation |  | 2020-11-04
v2 | VITO | First revision after partners' input |  | 2021-03-25
v2.1 | iMM | Revised version of template completed with pilot test case data |  | 2021-04-14

## Performance specifications
PRODUCT SPECIFICATION SHEET
Product name | Differential diagnostic PCR test (Pilot Test Case N°1)
Product reference | 0
Date | 2021-04-14 | IP rights | Instituto de Medicina Molecular João Lobo Antunes (iMM)
Key performance feature |  | Product status | Target values
APPLICATION ASPECTS
Intended use |  | Qualitative detection of RNA(-) in individuals that have been diagnosed with SARS-CoV-2 and that still test positive two weeks after symptom onset, although no longer showing clinical signs. This kit differentiates if the virus is still actively replicating (persistent infection) or if there is viral shedding, and therefore the virus is no longer active in samples from individuals of all ages.
Target use setting (central lab, triage center, home, etc) and volume setting (low, high) |  | Central lab, high-volume setting.
Target population/patient |  | Individuals of all ages that have been infected with SARS-CoV-2 and no longer show clinical signs two weeks after symptom onset, however still test positive for SARS-CoV-2 using traditional RT-PCR tests. | Individuals of all ages that have been infected with SARS-CoV-2 and no longer show clinical signs two weeks after symptom onset, however still test positive for SARS-CoV-2 using traditional RT-PCR tests. Proof of concept in oncological patients testing positive for SARS-CoV-2 by RT-PCR 20 days after the onset of symptoms, although meeting the clinical discharge criteria (according to DGS 004/2020 of 23/03/2020 updated on 14/10/2020) and that require a negative test to resume treatments (chemotherapy and/or immunotherapy)
Sample type |  | Nasopharyngeal/oropharyngeal sample.
Sample collection: compatible devices |  | Swabs in inactivation transport medium (kit can also be used in samples collected in non-inactivating media, although BSL3 facilities are required to inactivate swab samples before the use of this kit).
Target molecule/analyte |  | SARS-CoV-2 negative sense RNA
Detection principle |  | SARS-CoV-2 is a ssRNA(+) virus and its full-length genome is used as a template to produce a sub-genomic form of ssRNA(-) as a template for viral replication. As such, RNA(+) is present at much higher concentration in a sample and the detection of RNA(-) is an indicative of a replication-competent virus. If the virus is replication-competent, it can generate all the necessary components for a fully functional, and therefore infectious progeny. This kit uses a primer for SARS-CoV-2 RNA(-) and another for SARS-CoV-2 RNA(+) detection both with a non-SARS-CoV-2 tag sequence added in separate reactions during reverse transcription. The tag sequence is incorporated into the newly formed virus cDNA sequence. The same sample is tested in 2 replicates for the detection of RNA(+), and due to the low concentration of RNA(-), is tested in 10 replicates for RNA(-) detection. A fluorescently labeled probe and PCR primers are added to the mixture to hybridize with the target sequence (now containing the tag sequence) during PCR amplification. Fluorescence detection will only occur when the target sequence is present, resulting in a signal that is proportional to the amount of amplified product in the sample. Data is obtained as sigmoidal amplification curves of Ct value vs. cycle. Test results must be validated (see acceptance criteria) before data interpretation. If the sample contains infectious virus, then a positive result is considered when RNA(+) is detected in its corresponding duplicates and RNA(-) is detected in at least one of the 10 replicates. A negative result is considered when RNA(-) is not detected in any of the 10 replicates, even if RNA(+) is detected in two, one or none of its duplicates. An inconclusive result is considered when RNA(-) is detected in at least one of the 10 replicates, but RNA(+) is not detected or only detected in one of its duplicates. Ct values obtained in RT-PCR can only be considered positive and valid until the defined limit of detection obtained through a standard curve. | Additionally, when used as a multiplex RT-PCR, for a result to be considered positive, amplification of the RNAseP control gene must be observed in all sample reactions.
Test kit format (reagents, controls, consumables; sample collection, processing, analysis) |  | Kit provides controls and reagents with the exception of routine laboratory consumables and accessories for sample collection and transport. Reagents: A. Isolation using Viral RNA isolation kit from NZYtech (MB40701); B. Real-time RT-PCR using One-step NZY RT-PCR probe kit from NZYtech (MB351) containing reverse transcriptase and qPCR buffer with polymerase; one RT primer for RNA(+) and one RT primer for RNA(-); one set of PCR primers/probe specific for SARS-CoV-2 N genomic region. Consumables (not provided): RNAse-free eppendorfs, RNAse-free and aerosol-resistant tips, RNAse-free PCR strips, qPCR 96-well plates and adhesive foils. Sample collection (not provided): Nasopharyngeal/oropharyngeal swab samples collected in inactivation or non-inactivation transport medium. Processing and analysis: RNA extraction from swab samples, cDNA preparation, Real time PCR amplification, and result analysis using real-time RT-PCR software. | Controls to be included: RNA(+) and RNA(-) long oligos and RNAseP DNA.
End user profile (trained health care worker or lab technician, trained lay worker, patient, etc) |  | Lab technician or clinical laboratory personnel with training in techniques of real-time PCR. Sample collection should be performed by a trained healthcare provider only.
Training needs end user, duration (incl. sample collection, test procedure, data interpretation, QC and biosafety) |  | End user needs to be trained on working under sterile conditions, RNA isolation and real-time PCR procedures, and interpretation of real-time RT-PCR data.
QUALITY CONTROL
Positive control material/method |  | Low Ct sample from a SARS-CoV-2 infected individual tested using standard (non-tagged) real-time RT-PCR; or RNA isolated from viral culture supernatant obtained from inoculation of Vero CCL-81 cells with a SARS-CoV-2 isolate. | Synthetic RNA(+) and RNA(-) transcripts for the N gene of the SARS-CoV-2 genome.
Negative control material/method |  | NTC (no template control) with nuclease free water in place of sample.
Internal control material/method |  | None | Human RNAseP gene primer/probe set to show amplification in all clinical samples using multiplex real-time RT-PCR.
Calibration control method |  | RNA(+) standard curve obtained using a SARS-CoV-2 RNA standard with 200 copies/μL.
Type of data analysis (qualitative/quantitative) |  | Qualitative (can be quantitative for RNA(+) when using the standard curve obtained, if necessary).
Data interpretation type (visual manual read, digital hardware read, proprietary or smartphone app) |  | Ct curves and values obtained through real-time RT-PCR machine software; visual manual read on real-time RT-PCR software to determine positive and negative reactions.
Mathematical prediction model |  | None
Test acceptance criteria |  | A valid run must meet all the following criteria, otherwise the run is considered invalid and sample results cannot be interpreted: 1) Detection of RNA(+) in both duplicates of the positive control; 2) Detection of RNA(-) in all 10 replicates of the positive control; 3) Amplification of RNA(+) 9-12 Cts lower than RNA(-); and 4) no amplification on NTC control, assuring the absence of contamination during assay set-up. | Additionally, when used as a multiplex RT-PCR, the following validation criteria are added: 5) Detection of RNAseP control gene in all 12 reactions of the positive control; and 6) RnaseP detection in all extracted samples to monitor RNA extraction and reverse transcription.
ANALYTICAL PERFORMANCE
Sensitivity/Limit of detection |  | Using a SARS-CoV-2 RNA(+) standard, one copy/reaction was detected at a Ct value of 36 for RNA(+); In a viral culture sample, RNA(-) was detected until a Ct of 37, which corresponded to a Ct of 29 for RNA(+). | Use of a SARS-CoV-2 RNA(-) standard to determine corresponding Ct value for one copy/reaction.
Specificity |  | Primers and probes used were designed by CDC to amplify a specific sequence region of the nucleocapsid (N) gene of the SARS-CoV-2 genome. Additionally, a non-SARS-CoV-2 sequence was added to the RT primer. In silico analysis of the primers/probe set using NCBI BLAST discarded any cross-reactivity with other SARS viruses and commonly related pathogens (e.g. bacteria and other respiratory viruses). Primers/probes (excluding tag primer) were checked by CDC for false-positive results, as previously described: Lu, X., Wang, L., Sakthivel, S. K., Whitaker, B., Murray, J., Kamili, S....Lindstrom, S. (2020). US CDC Real-Time Reverse Transcription PCR Panel for Detection of Severe Acute Respiratory Syndrome Coronavirus 2. Emerging Infectious Diseases, 26(8), 1654-1665. https://dx.doi.org/10.3201/eid2608.201246. | Specificity of primers and probes will be validated in vitro to test for potential cross-reactivity with a panel of common respiratory pathogens (e.g. bacteria and other respiratory viruses).
PCR efficiency |  | 89,78% for RNA(+) and 82,12% for RNA(-).
Robustness |  | RNA(-) was detected using different RNA extraction methods (manual or automated), different RT-PCR machines, and different positive controls (positive samples or viral culture samples).
Precision |  | Each sample is analyzed in duplicate for RNA(+) and in 10 replicates for RNA(-).
Analysis of cross reactivity |  | In silico analysis of the primers/probe set using NCBI BLAST discarded any cross-reactivity with other SARS viruses and commonly related pathogens (e.g. bacteria and other respiratory viruses). Primers/probes (excluding tag primer) were designed by CDC and checked for false-positive results, as previously described: Lu, X., Wang, L., Sakthivel, S. K., Whitaker, B., Murray, J., Kamili, S....Lindstrom, S. (2020). US CDC Real-Time Reverse Transcription PCR Panel for Detection of Severe Acute Respiratory Syndrome Coronavirus 2. Emerging Infectious Diseases, 26(8), 1654-1665. https://dx.doi.org/10.3201/eid2608.201246. | Specificity of primers and probes will be validated in vitro to test for potential cross-reactivity with a panel of common respiratory pathogens (e.g. bacteria and other respiratory viruses).
Analysis of interference |  | Each run contained a negative control consisting of a pool of negative samples extracted and analyzed in the same conditions as the positive samples. | Inclusion of RNAseP primer/probe set in kit to amplify at a determined Ct value in every multiplex reaction as an endogenous internal control will demonstrate that polymerase is not inhibited and therefore there are no interfering substances.
Reproducibility |  | RNA obtained from a viral culture sample was diluted until the determined limit of detection for RNA(-). This dilution was reverse transcribed six times and each RT reaction was amplified in duplicate for RNA(+) and in 10 replicates for RNA(-). RNA(-) was consistently detected in at least one of the 10 replicates.
CLINICAL PERFORMANCE
Sample origin |  | Portuguese police force and hospital patients. | Oncological patients testing positive for SARS-CoV-2 20 days after the onset of symptoms.
Study design (prospective/retrospective sampling, cohort size, inclusion/exclusion criteria etc.) |  | Retrospective study with 8 individuals from the Portuguese police force and 5 hospital patients. | Prospective unicentric study with an ideal size of 30 oncological patients.
Ethical approval obtained, yes/no |  | No, but samples were collected with informed consent. | Awaiting ethical approval.
Sensitivity/Limit of detection |  | As in analytical performance. | To be confirmed.
Specificity |  | As in analytical performance. | To be confirmed.
PCR efficiency |  | As in analytical performance. | To be confirmed.
Robustness |  | As in analytical performance. | To be confirmed.
Precision |  | As in analytical performance. | To be confirmed.
Analysis of cross reactivity |  | As in analytical performance. | To be confirmed.
Analysis of interference |  | As in analytical performance. | To be confirmed.
Reproducibility |  | As in analytical performance. | To be confirmed.
TEST PROCEDURE
Sample collection reagents |  | Swabs with inactivation transport medium.
Sample storage and stability time pre-testing |  | Swabs can be stored at 2-8°C for up to 72 hours after collection. Long-term storage can be done at -80ºC if needed.
Sample minimal volume |  | 2 mL
Need to transfer a precise volume of sample |  | Yes
Sample preparation steps: number |  | 5 (sample heat-inactivation, RNA isolation, cDNA preparation, RT, real-time PCR).
Test kit stability and storage conditions |  | Reagents from One-step NZY RT-PCR probe kit from NZYtech (MB351) to be stored at -20 °C in a constant temperature freezer. Minimize exposure of the RT-PCR mix to room temperature and avoid over exposure of the primer-probe mixes to light. Do not leave components out at room temperature for long periods.
Reagents reconstitution |  | Not needed
Stability of kit once opened |  | Not determined
Specimen capacity and throughput |  | 5 samples + 2 control samples (positive and negative) per 96-well plate. One real-time RT-PCR kit is enough for 10 plates.
Safety precautions, sample collection to test result (biosafety, chemical safety, protective equipment) |  | While testing the samples, a facemask, lab coat and gloves need to be used at all times. A sterile environment needs to maintained, samples should be handled/prepared in a separate area from where PCR reagents are handled.
Operating conditions (temperature, relative humidity, altitude) |  | Room temperature
Number of timed steps (use of different reagents, incubation steps) |  | 2 (Sample heat-inactivation for 30 minutes before RNA isolation; 3-minute incubation in column membrane during RNA elution in RNA isolation protocol)
Time to result |  | 3h30 (from sample heat inactivation to result) | Can be improved through use of automated RNA isolation and RNA sample distribution in 96-well plates for real-time RT-PCR
Result validity stability |  | End-point result
Invalid rate with correct use by operator |  | Not determined
Remote connectivity capacity for data capture |  | Data can be uploaded to a cloud and accessed through the real-time RT-PCR software in every computer
Additional equipment: compatible readers, battery |  | NA
Need for maintenance/spare parts |  | Maintenance of thermal cycler and RT-PCR machine.
Waste/disposal requirements |  | Liquids should be collected in leak-proof containers such as flasks or bottles. Solid laboratory waste should be collected in an appropriate waste bag. Waste should be treated as biological waste.
Accessibility (production scale, pricing, geographical area)
Regulatory compliance
Target MD/IVD classification |  | IVDR Rule 1 Class D. IVDD 'All others'
Quality management system, risk management systems employed |  | Available data indicates that adults with mild to moderate COVID-19 remain infectious no longer than 10 days after symptom onset. Most adults with more severe to critical illness or severe immunocompromise likely remain infectious no longer than 20 days after symptom onset; however, there have been several reports of people shedding replication-competent virus beyond 20 days due to severe immunocompromise. The traditional diagnosis RT-PCR tests cannot determine whether the person being tested is emitting whole virus capable of infecting someone else, or if the swabs instead are collecting the viral debris that can be shed after an infection. Currently, to determine potential virus infectiousness in a sample collected in non-inactivating media, cells have to be infected with swab samples under BSL3 conditions. This process is time-consuming, requires specialized training and not all infrastructures have access to BSL3 facilities.This kit is a RT-PCR test for the qualitative detection of the viral negative-sense RNA from SARS-CoV-2 in nasopharyngeal and/or oropharyngeal specimens. It offers the possibility to distinguish COVID-19 patients with replicating virus from those with a non-replicating status without the need of BSL3 conditions. This kit should be used in individuals that have been diagnosed with SARS-CoV-2 and that still test positive two weeks after symptom onset, although no longer showing clinical signs allowing for better medical decisions when discharging patients. It should not be applied as a screening method to the general population to replace traditional RT-PCR diagnostic tests, as even though it offers unique and specific information on the infectivity status of the virus, is more laborious and traditional diagnosis tests are enough to correctly diagnose individuals in the first two weeks after symptom onset when the viral load is high.
Status of regulatory approvals (e.g. CE/FDA)